Technical specification:

- General requirements for the full refrigerator system (dry dilution refrigerator with automated gas handling system, control electronics, temperature control system, wiring and superconducting magnet):

- Ready to use brand new mass-produced (not prototype) cryogen free dilution refrigerator (3He/4He)-based system with superconducting magnet without a need to fill with liquid He.

- 3He/4He mixture suitable to run the system.

- Fully automated operation involving cool-down process from room temperature to base temperature and stabilizing the system temperature at given setpoint.

- Both Automatic and Manual control possible.

- Monitoring and controlling of the system through network possible.

- Fully modular system – dilution unit, pulse tube refrigerator, magnet, all wiring and all services must be field serviceable and field replaceable with no desoldering or breaking of welded joints.

- Advanced vibration insulation for cryocooler (should be no rigid connection of pulse tube to any stage of the dry dilution refrigerator including top flange of the refrigerator). The turbo pump should be connected to the refrigerator in the way damping mechanical vibrations.

- A standing anti-vibration frame for the refrigerator system.

- Dedicated magnet providing magnetic field minimum 12T.

- Reliable electric connections.

- Reliable connections of the system circulating the helium mixture

- External LN2 trap.

- The whole system should fit to the lab space prepared for the system: laboratory 1: height 3.1m area 6m x 7m laboratory 2: height 3.1m area 6m x 2.5m

- Detailed specification of the refrigerator:

- Base temperature 12 mK with electrical lines described in p.4. and with a magnet, running with 13L of 3He/4He mixture.

- Cooling power at sample stage (at 100 mK, measured in vacuum on base of mixing chamber (not in liquid or inside the mixing chamber)): 300 W or more.

- Free cooling power at 4K stage at least 200mW.

- Cooling power (at 20 mK, measured in vacuum on base of mixing chamber (not in liquid or inside the mixing chamber)) at sample stage: 10 W or more.

- Pulse tube cryocooler with cooling power high enough to support 12T magnet. The pulse tube cooling power should be large enough for an upgraded system with magnet providing larger magnetic field.

- Cool-down time to base temperature (12mK or below) without magnet 30 hours or less from start of the PT.

- Cool-down time to base temperature (12mK or below) with magnet 70 hours or less.

- Refrigerator designed and installed in a manner providing with a vibration-free operation – anti-vibration bellows in frame, isolation for Turbo Pump and PTR.

- Construction allowing for installation of a dedicated magnet. Full set of components (radiation and vacuum shields etc.) necessary for operation with and without magnet should be included.

- Oil free pumps must be supplied with the system. All pumps must operate within original manufacturers specifications in all modes of operation – condensing, circulating.

- Sample space – at least 360mm diameter x 250mm length, in vacuum.

- System must be able to operate with continuous temperature control from base temperature to minimum 30K with integrated magnet at full field and temperature stabilization with an accuracy of 1mK at temperatures T<100mK and 1% at T>100mK.

- Detailed specification for the solenoid magnet integrated in the dry dilution refrigerator:

- Superconducting magnet with persistence mode switch and power supplies fully integrated and dedicated to the dry dilution refrigerator.

- Magnetic field along Z-axis ranging from 0 till 12T or more .

- Magnet power supply with both manual and automatic control.

- Clamp connections for magnet current leads allowing for easy electrical connection and disconnection of the magnet without a need for soldering

- Possibility to uninstall magnet effortlessly and run the dry dilution refrigerator without magnet,

- Dilution refrigerator and integrated superconducting magnet should have fully automated functionality. This control software should allow remote access from a data collection computer with full automation, status read and set point write control functions incorporated.

- Field uniformity inside magnet should be not worse than 0.1% total homogeneity over 1cm in DSV for Z-direction.

- Magnet bore size min.77 mm.

- Magnet fully protected against damage due to quench or power cuts, including integrated diode quench protection circuit, quench detection and auto run-down functionality.

- Requirements for the electrical lines:

- Experimental DC leads to the mixing chamber:

- 24 copper twisted pairs dc wires (12 pairs) from room temperature flange to 4K flange 24 NbTi twisted pairs dc wires (12 pairs) from 4K plate to MXC

- Experimental RF leads to the mixing chamber (The RF wires should be a loop-shaped to replace them easily.):

- 4 x CuNi-CuNi attenuated semi-rigid co-axial line from room temperature connector box to the mixing chamber flange.

- 2x CuNi-CuNi attenuated semi-rigid co-axial line from room temperature connector box to the 4K flange. 2x NbTi-NbTi non-attenuated semi-rigid superconducting co-axial line from 4K flange to MXC

- 12 x S1 st.st flexible coax from room temperature flange to mixing chamber.

- Low resistance wiring for an Attocube rotator 24-way DC Nomex piezo-electric nano-positioner.

- One low-noise amplifier mounted at 4K-plate including power supply block, HEMT supply, bracket and power cables and necessary DC loom.

- One low temperature circulator mounted at 4K-plate including bracket.

- Precisely 10mK calibrated RuOx thermometer (generic calibration unacceptable) for measurements of a sample on the mixing chamber.

Note: All lines must be properly thermalized at all stages of the Dilution Refrigerator.

- A separate OVC turbo pump.

- Four off optical fibers (multi mode) from the room temperature flange down to MXC.

- Additional features:

- Clamp connection for magnet current leads – six bolted electrical joints for a coil mounted on 4K plate at top of magnet for ease of magnet exchange (already described in p.3.4).

- Galvanic isolation of the dry dilution refrigerator from the frame and pump system.

- Manuals, spares kit and toolbox necessary to run the system.

- Software for controlling the system

System software should automatically log all system activities. This includes valves and pumps status, all pressure gauges, turbo pump speed, thermometry, magnet current and field.

The software automatically detects an unexpectedly high magnet temperature and halts an ongoing sweep of the magnet to minimize the risk of quenching.

System should allow remote operation through a standard internet protocol. Remote operation allows the user to control the refrigerator as well as the magnet, read all temperatures and change magnet and temperature PID set points and control all heaters.

All thermometers are read out using a 4-wire configuration (no common grounds).

In-line relief valves to protect the system in the event of a long term power cut.

- The manufacturer will conduct training on the use and necessary maintenance of the system for at least 3 users

- Delivery, installation and training by 31st of December 2023 year

- Two years fully comprehensive warranty .